**Performance characteristics of five antigen-detecting rapid diagnostic test (Ag-RDT) for SARS-CoV-2 asymptomatic infection: a head-to-head benchmark comparison**

**Supplementary materials**

**Table S1. Main characteristics of five rapid SARS-CoV-2 rapid antigen tests (according to the manufacturer)**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Assay name** | PanBioTM COVID-19 Ag Rapid Test Device | CLINITEST®Rapid COVID-19 Antigen Test | SARS-CoV-2 Rapid Antigen Test | SARS-CoV-2 Antigen Rapid Test Kit | COVID-19 Antigen Test Cassette  |
| **Distributor** | Abbott RapidDiagnostics | Siemens Healthineers (Shangai International Holding Corp) | Roche Diagnostics | Lepu Medical (Europe) cooperatief U.A. | SureScreen Diagnostics |
| **Manufacturer** | Panbio Ltd. | Healgen Scientific Limited Liability Company | SD Biosensor | Beijing Lepu Medical Technology Co., Ltd. | SureScreen Diagnostics Ltd |
| **Country of origin** | USA | USA | Republic of Korea | P.R. China | UK |
| **Certification** | CE-IVD | CE-IVD | CE-IVD | CE-IVD | CE-IVD |
| **Specimen** | Nasopharyngeal swab | Nasopharyngeal or nasal swab | Nasopharyngeal swab | Nasopharyngeal or oropharyngeal swab | Nasopharyngeal or oropharyngeal swab |
| **Intended use**  | Suspected cases of COVID-19 (clinical and/or epidemiological criteria) | Suspected cases of COVID-19 | Suspected cases of COVID-19 | Clinical samples | Suspected cases of COVID-19 within first 2weeks of symptoms |
| **Limit of detection****(manufacturer)** | 2,5 X101,8 TCID50/mL | 1,2 x 102 TCID50/mL | 3,1 X 102,2 TCID50/mL | Not specified | 2x10 2,4TCID50/mL |
| **Reported sensitivity and specificity** | 97.9% (95.2-99.3%)99.4% (97.0-100.0%) | 96.7% (88.7-99.60%)99.22% (97.21-99.91%) | 96.52% (91.33-99.04%)99.68% (98.22-99.99%) | No less than 90%No less than 100% | 94.55% (83.93-98.58%)100% (97.1-100%) |
| **Format** | Cassette | Cassette | Cassette | Card | Cassette |
| **Volume in extraction tube** | Marked line (~300uL) | 10 drops | Aliquoted (~250uL) | NA | 10 drops |
| **Volume applied****into cassette** | 5 drops | 4 drops | 3 drops | 6 drops | 2 drops |
| **Incubation** | 15-20 minutes | 15-20 minutes | 15-30 minutes | 15-20 minutes | 15 minutes |
| **Readout** | Visual: colored band | Visual: colored band | Visual: colored band | Visual: colored band | Visual: colored band |
| **UTM protocol** | Not specified | Not specified | Reported | Not specified | Not specified  |
| **Cross-reactivity against other human virus** | Reported | Reported | Reported | Reported | Reported  |

**Table S2.** Discordance analysis

|  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|   |   |   |   |   |   |   |   | **Cases** |
| **Abbott** | **Siemens** | **Roche** | **Lepu** | **SureScreen** | **All samples** | **PCR+**  | **%PCR+** | **<30** | **30-35** | **>35** |
| **Full agreement** |
| + | + | + | + | + | 29 | 29 | 100 | 21 | 7 | 1 |
| - | - | - | - | - | 193 | 37 | 19.17 | 3 | 20 | 14 |
| **Discordance** |
| - | - | - | + | - | 23 | 8 | 34.78 | 1 | 6 | 1 |
| - | + | - | - | - | 10 | 8 | 80 |  | 5 | 3 |
| - | - | + | - | - | 4 | 2 | 50 |  |  | 2 |
| + | - | - | - | - | 1 | 1 | 100 |  |  | 1 |
| - | - | - | - | + | 4 | 0 | 0 |  |  |  |
| - | + | - | + | - | 3 | 3 | 100 | 1 | 2 |  |
| - | + | + | - | - | 4 | 3 | 75 | 1 | 2 |  |
| - | - | + | + | - | 4 | 0 | 0 |  |  |  |
| + | - | - | + | - | 1 | 0 | 0 |  |  |  |
| - | + | + | + | - | 1 | 1 | 100 | 1 |  |  |
| + | + | + | - | - | 3 | 3 | 100 | 1 | 1 | 1 |
| + | + | + | + | - | 5 | 5 | 100 | 1 | 2 | 2 |
| + | - | + | - | - | 1 | 1 | 100 |   | 1 |   |

**Table S3.** Estimated false negative and false positives from the positive- and negative-predictive values modeled for each prevalence.

|  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | **All samples** | **Samples CT<30** | **Samples CT [30-35]** | **Samples CT>35** |
|  | **Prevalence** | **Specificity** | **Sensitivity** | **FN** | **FP** | **Total****Errors** | **PPV** | **NPV** | **PPV** | **NPV** | **PPV** | **NPV** | **PPV** | **NPV** |
| Abbott | 0.01 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 6 (5-7) | 5 (0-30) | 12 (5-37) | 41.91  (8.99-97.3) | 99.38 (99.27-99.49) | 58.89 (16.39-98.52) | 99.76 (99.56-99.9) | 30.88 (4.1-96.62) | 99.23 (99.1-99.39) | 27.21 (2.27-96.78) | 99.19 (99.04-99.4) |
| Siemens | 0.01 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 5(4-6) | 16 (3-47) | 21 (7-53) | 24.28 (8.21-64.94) | 99.5(99.38-99.61) | 35.06 (13.04-74.33) | 99.86 (99.68-99.96) | 20.46 (5.52-63.08) | 99.4  (99.23-99.56) | 14.85 (2.55-59.78) | 99.27 (99.08-99.49) |
| Roche | 0.01 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 6 (5-7) | 38(16-77) | 44 (20-84) | 10.42 (4.27-26.15) | 99.41 (99.28-99.53) | 18.2 (7.94-38.31) | 99.83 (99.62-99.94) | 7.02 (2.07-22.24) | 99.25  (99.09-99.42) | 6.02 (1.22-22.9) | 99.21 (99.02-99.44) |
| Lepu | 0.01 | 89.19%  (83.8-93.27) | 45.54% (35.6-55.76) | 5 (4-6) | 109 (68-164) | 115 (72-170) | 4.08 (2.17-7.72) | 99.39 (99.23-99.52) | 7.22 (3.91-12.41) | 99.81 (99.58-99.94) | 3.34 (1.43-7.3) | 99.29  (99.08-99.49) | 1.47 (0.28-5.14) | 99.06 (98.86-99.31) |
| Surescreen | 0.01 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 7 (6-8) | 22 (6-55) | 29 (12-63) | 11.83(3.6-39.68) | 99.27 (99.15-99.38) | 24.64 (8.58-59.26) | 99.69 (99.48-99.85) | 6.64 (1.16-32.99) | 99.13  (99.01-99.28) | 1.83 (0.02-25.77) | 99.02 (98.94-99.2) |
| Abbott | 0.05 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 31 (26-35) | 6 (0-31) | 36 (26-67) | 78.99 (33.98-99.47) | 96.85 (96.3-97.38) | 88.19 (50.52-99.71) | 98.78 (97.76-99.48) | 69.96 (18.21-99.33) | 96.13  (95.47-96.88) | 66.07 (10.78-99.37) | 95.94(95.19-96.97) |
| Siemens | 0.05 | 98.38%(95.33-99.66) | 51.49% (41.33-61.55) | 24 (19-29) | 17 (4-49) | 41 (23-78) | 62.56 (31.8-90.61) | 97.47 (96.86-98.01) | 73.77 (43.87-93.79) | 99.29(98.33-99.8) | 57.28 (23.34-89.9) | 96.96  (96.13-97.77) | 47.61 (11.99-88.56) | 96.29 (95.37-97.4) |
| Roche | 0.05 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 28 (23-33) | 40 (16-80) | 68(39-114) | 37.73 (18.85-64.85) | 97.01 (96.36-97.59) | 53.69 (31.02-76.39) | 99.1 (98.06-99.7) | 28.22 (9.92-59.85) | 96.22 (95.43-97.07) | 25.03 (6.06-60.75) | 96.01 (95.09-97.15) |
| Lepu | 0.05 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 27 (22-32) | 114 (71-171) | 141 (93-203) | 18.15 (10.37-30.36) | 96.89 (96.11-97.56) | 28.86 (17.5-42.46) | 99.03 (97.87-99.68) | 15.25 (7.01-29.09) | 96.41 (95.4-97.39) | 7.23 (1.45-22.01) | 95.28 (94.34-96.52) |
| Surescreen | 0.05 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 36 (31-40) | 23 (6-57) | 58 (37-97) | 41.14 (16.3-77.42) | 96.31 (95.74-96.85) | 63.02 (32.85-88.34) | 98.41 (97.32-99.23) | 27.03 (5.78-71.96) | 95.64  (95.05-96.37) | 8.87 (0.1-64.4) | 95.09 (94.73-95.95) |
| Abbott | 0.1 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 61 (51-71) | 6 (0-33) | 67 (51-104) | 88.81 (52.07-99.75) | 93.58 (92.49-94.62) | 94.03 (68.31-99.86) | 97.46 (95.38-98.91) | 83.1 (31.98-99.68) | 92.17 (90.9-93.63) | 80.43 (20.33-99.7) | 91.8 (90.36-93.82) |
| Siemens | 0.1 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 49 (38-59) | 18 (4-52) | 67(42-111) | 77.91 (49.6-95.32) | 94.81 (93.6-95.89) | 85.59 (62.26-96.96) | 98.52 (96.54-99.58) | 73.89 (39.13-94.95) | 93.78 (92.16-95.4) | 65.74 (22.33-94.24) | 92.48 (90.7-94.66) |
| Roche | 0.1 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 56 (46-66) | 42 (17-85) | 98 (63-151) | 56.13 (32.9-79.57) | 93.88 (92.62-95.04) | 70.99 (48.7-87.23) | 98.11 (95.99-99.37) | 45.35 (18.86-75.88) | 92.35 (90.82-94) | 41.34 (11.98-76.57) | 91.93 (90.17-94.17) |
| Lepu | 0.1 | 89.19%(83.8-93.27) | 45.54% (35.6-55.76) | 54 (44-64) | 120 (75-180) | 175 (119-244) | 31.88 (19.62-47.93) | 93.65 (92.13-94.99) | 46.13 (30.92-60.91) | 97.97 (95.6-99.33) | 27.53 (13.73-46.41) | 92.72  (90.76-94.64) | 14.12 (3.02-37.33) | 90.53 (88.76-92.92) |
| Surescreen | 0.1 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 71 (61-80) | 24 (7-60) | 95 (68-140) | 59.6 (29.14-87.86) | 92.51 (91.42-93.57) | 78.25(50.81-94.12) | 96.71 (94.51-98.38) | 43.88 (11.47-84.42) | 91.22  (90.09-92.64) | 17.05 (0.21-79.25) | 90.17 (89.49-91.83) |
| Abbott | 0.15 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 92 (77-106) | 6(0-35) | 98 (77-141) | 92.65 (63.31-99.84) | 90.18 (88.58-91.72) | 96.16 (77.4-99.91) | 96.02(92.86-98.28) | 88.65 (42.75-99.8) | 88.11  (86.28-90.25) | 86.72 (28.84-99.81) | 87.57 (85.51-90.53) |
| Siemens | 0.15 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 73 (58-88) | 19(4-55) | 92 (62-143) | 84.85 (60.99-97) | 91.99 (90.2-93.63) | 90.41 (72.38-98.06) | 97.66 (94.62-99.34) | 81.8 (50.52-96.76) | 90.47 (88.1-92.89) | 75.29 (31.35-96.29) | 88.56 (86-91.78) |
| Roche | 0.15 | 96.22%(92.36-98.47) | 43.56% (33.72-53.8) | 85 (69-99) | 45(18-90) | 129 (87-189) | 67.02 (43.78-86.08) | 90.62 (88.76-92.35) | 79.54 (60.12-91.56) | 97.03 (93.78-99) | 56.86 (26.97-83.33) | 88.37 (86.17-90.8) | 52.82 (17.77-83.84) | 87.77 (85.24-91.05) |
| Lepu | 0.15 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 82 (66-97) | 127 (79-191) | 209 (146-287) | 42.64 (27.94-59.38) | 90.27 (88.06-92.28) | 57.63 (41.56-71.22) | 96.81 (93.19-98.94) | 37.63 (20.18-57.9) | 88.91  (86.08-91.75) | 20.71 (4.71-48.62) | 85.75 (83.26-89.21) |
| Surescreen | 0.15 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 107 (92-120) | 25 (7-64) | 132 (99-184) | 70.09 (39.51-92) | 88.61 (87.03-90.17) | 85.1 (62.13-96.21) | 94.87 (91.56-97.45) | 55.4 (17.06-89.59) | 86.74  (85.12-88.79) | 24.61 (0.33-85.85) | 85.24 (84.29-87.61) |
| Abbott | 0.2 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 123 (102-142) | 7 (0-37) | 130 (103-179) | 94.7 (70.97-99.89) | 86.63 (84.55-88.66) | 97.26 (82.91-99.94) | 94.46 (90.18-97.58) | 91.71 (51.4-99.86) | 83.95  (81.62-86.72) | 90.24 (36.47-99.87) | 83.26 (80.64-87.09) |
| Siemens | 0.2 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 97 (77-117) | 20(4-58) | 117 (81-176) | 88.81 (68.89-97.87) | 89.02 (86.67-91.2) | 93.04 (78.78-98.62) | 96.72 (92.54-99.07) | 86.43 (59.13-97.69) | 87.02  (83.93-90.22) | 81.19 (39.28-97.35) | 84.53 (81.26-88.73) |
| Roche | 0.2 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 113(92-133) | 47 (19-96) | 160 (112-228) | 74.22 (52.46-89.76) | 87.21 (84.79-89.5) | 84.63 (68.11-93.89) | 95.85 (91.41-98.59) | 65.12 (34.34-87.62) | 84.29  (81.47-87.45) | 61.33 (23.44-88.03) | 83.51 (80.3-87.77) |
| Lepu | 0.2 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 109 (88-129) | 135 (84-203) | 244 (173-331) | 51.3 (35.46-67.44) | 86.76 (83.88-89.4) | 65.84 (50.18-77.8) | 95.54 (90.61-98.51) | 46.08(26.37-66.08) | 84.98 (81.36-88.7) | 27.01 (6.54-57.27) | 80.94 (77.83-85.37) |
| Surescreen | 0.2 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 143 (123-160) | 27 (7-68) | 170 (130-228) | 76.85 (48.06-94.21) | 84.59 (82.57-86.62) | 89 (69.92-97.3) | 92.88 (88.45-96.43) | 63.76 (22.56-92.42) | 82.19  (80.15-84.83) | 31.62 (0.46-89.57) | 80.3 (79.11-83.31) |
| Abbott | 0.5 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 307 (256-355) | 11 (0-59) | 318 (256-414) | 98.62 (90.72-99.97) | 61.84 (57.78-66.14) | 99.3 (95.1-99.98) | 81 (69.65-90.96) | 97.79 (80.88-99.96) | 56.66  (52.61-62.02) | 97.37 (69.66-99.97) | 55.42 (51.01-62.77) |
| Siemens | 0.5 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 243 (192-293) | 32 (7-93) | 275 (199-387) | 96.95 (89.86-99.46) | 66.97 (61.9-72.16) | 98.16(93.69-99.65) | 88.06 (75.63-96.37) | 96.22 (85.26-99.41) | 62.63  (56.63-69.75) | 94.53 (72.12-99.32) | 57.74 (52.02-66.32) |
| Roche | 0.5 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 282 (231-331) | 76 (31-153) | 358 (262-484) | 92.01 (81.53-97.23) | 63.03 (58.22-68.06) | 95.66 (89.52-98.4) | 85.24(72.68-94.58) | 88.19 (67.66-96.59) | 57.29  (52.37-63.52) | 86.38 (55.05-96.71) | 55.87 (50.47-64.22) |
| Lepu | 0.5 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 272 (221-322) | 216(135-324) | 488 (356-646) | 80.82 (68.72-89.23) | 62.09 (56.54-67.83) | 88.52 (80.12-93.34) | 84.26 (70.7-94.3) | 77.37 (58.89-88.63) | 58.59  (52.18-66.23) | 59.68 (21.88-84.28) | 51.5 (46.75-59.34) |
| Surescreen | 0.5 | 97.84%(94.56-99.41) | 28.71% (20.15-38.57) | 356 (307-399) | 43 (12-109) | 400 (319-508) | 93(78.73-98.49) | 57.85 (54.21-61.81) | 97 (90.29-99.31) | 76.53 (65.69-87.09) | 87.56 (53.82-97.99) | 53.57  (50.24-58.29) | 64.91 (1.83-97.17) | 50.47 (48.63-55.52) |
| Abbott | 0.85 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 522 (435-603) | 36 (1-198) | 558 (436-801) | 99.75 (98.23-100) | 22.23 (19.45-25.64) | 99.88 (99.1-100) | 42.93 (28.82-63.98) | 99.6 (96-99.99) | 18.74  (16.38-22.37) | 99.53 (92.86-99.99) | 17.99 (15.52-22.93) |
| Siemens | 0.85 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 412 (327-499) | 108(22-311) | 520 (349-810) | 99.45 (98.05-99.9) | 26.35(22.28-31.39) | 99.67(98.83-99.94) | 56.56 (35.38-82.4) | 99.31 (97.04-99.9) | 22.83  (18.73-28.92) | 98.99 (93.61-99.88) | 19.43 (16.06-25.79) |
| Roche | 0.85 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 480 (393-563) | 252(102-509) | 732 (495-1073) | 98.49 (96.16-99.5) | 23.13 (19.74-27.33) | 99.21 (97.98-99.71) | 50.46 (31.95-75.49) | 97.69 (92.22-99.38) | 19.14  (16.25-23.51) | 97.29 (87.4-99.4) | 18.26 (15.24-24.05) |
| Lepu | 0.85 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 463 (376-547) | 721 (449-1080) | 1184 (825-1628) | 95.98 (92.57-97.91) | 22.42 (18.67-27.12) | 97.76 (95.8-98.76) | 48.57 (29.87-74.47) | 95.09 (89.03-97.79) | 19.98  (16.15-25.71) | 89.35 (61.35-96.81) | 15.78 (13.41-20.48) |
| Surescreen | 0.85 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 606 (522-679) | 144 (39-363) | 750(562-1042) | 98.69 (95.45-99.73) | 19.5 (17.28-22.21) | 99.46 (98.14-99.88) | 36.53 (25.25-54.35) | 97.55 (86.85-99.64) | 16.92  (15.12-19.78) | 91.29 (9.53-99.49) | 15.24 (14.31-18.05) |
| Abbott | 0.9 | 99.46% (97.03-99.99) | 38.61% (29.09-48.82) | 552(461-638) | 54 (1-297) | 607 (462-936) | 99.84 (98.88-100) | 15.26 (13.2-17.84) | 99.92(99.43-100) | 32.14 (20.32-52.79) | 99.75 (97.44-100) | 12.68  (10.98-15.36) | 99.7 (95.38-100) | 12.14 (10.37-15.78) |
| Siemens | 0.9 | 98.38% (95.33-99.66) | 51.49% (41.33-61.55) | 437 (346-528) | 162 (34-467) | 599 (380-995) | 99.65 (98.76-99.94) | 18.39 (15.29-22.36) | 99.79 (99.26-99.96) | 45.05 (25.64-74.68) | 99.57(98.12-99.93) | 15.7  (12.67-20.39) | 99.36 (95.88-99.92) | 13.18 (10.75-17.95) |
| Roche | 0.9 | 96.22% (92.36-98.47) | 43.56% (33.72-53.8) | 508 (416-597) | 378 (153-764) | 886 (569-1361) | 99.04 (97.54-99.68) | 15.93 (13.41-19.15) | 99.5 (98.72-99.82) | 39.08 (22.81-65.98) | 98.53 (94.96-99.61) | 12.97  (10.89-16.21) | 98.28 (91.68-99.62) | 12.33 (10.17-16.62) |
| Lepu | 0.9 | 89.19% (83.8-93.27) | 45.54% (35.6-55.76) | 490 (398-580) | 1081 (673-1620) | 1571 (1071-2200) | 97.43 (95.19-98.68) | 15.4 (12.63-18.98) | 98.58 (97.32-99.21) | 37.29 (21.15-64.75) | 96.85 (92.8-98.59) | 13.58 (10.81-17.9) | 93.02 (71.6-97.97) | 10.55 (8.89-13.95) |
| Surescreen | 0.9 | 97.84% (94.56-99.41) | 28.71% (20.15-38.57) | 642 (553-719) | 216 (59-544) | 858 (612-1263) | 99.17 (97.09-99.83) | 13.23 (11.63-15.24) | 99.66 (98.82-99.92) | 26.6 (17.54-42.84) | 98.45 (91.3-99.77) | 11.36  (10.09-13.44) | 94.33 (14.34-99.68) | 10.17 (9.52-12.18) |