Supplementary Table S3. Intra- and inter-assay variation of FMIA for each antibody class and antigen.

|  |  |  |  |
| --- | --- | --- | --- |
| Antibody class | Antigen | Intra-assay variation (CV%)a | Inter-assay variation (CV%) |
| Within a dayb | Between daysc |
| IgG | N | 7 | 6 | 15 |
| RBD | 8 | 4 | 11 |
| SFL | 8 | 2 | 11 |
| Mean | 8 | 4 | 12 |
| IgA | N | 7 | 2 | 6 |
| RBD | 13 | 9 | 13 |
| SFL | 13 | 20 | 7 |
| Mean | 11 | 10 | 9 |
| IgM | N | 10 | 3 | 12 |
| RBD | 10 | 4 | 7 |
| SFL | 9 | 5 | 7 |
| Mean | 10 | 4 | 9 |

CV = coefficient of variation, N = SARS-CoV-2 (Wuhan-Hu-1) nucleoprotein, RBD = receptor binding domain of SARS-CoV-2 (Wuhan-Hu-1) spike glycoprotein, SFL = full length spike glycoprotein of SARS-CoV-2 (Wuhan-Hu-1), a = Calculated from antibody concentrations of in-house standards that were analysed in 13 assays as seven serial dilutions (1:400–1:1638400) each with two replicates. Results are the mean CV% of replicates within plates. b = In-house standard’s variation within one day. The mean of two replicates was used in the calculation of CV% between plates. Depending on the antibody class, values are based on 5–8 days and 10–17 assays (IgM: 5 days and 10 assays, IgA: 6 days and 12 assays, IgG: 8 days and 17 assays). c = Five sera analysed on five days as two replicates. The mean of two replicates was used in calculations of CV% between days.