**Supplementary Table 1**: Results of biological indicator (BI) monitoring after one-cycle of sterilization in a SteraMist-equipped environment chamber

|  |  |  |  |
| --- | --- | --- | --- |
| **Representative PPE items and BI location** | **Was surface pre-treated?** | **7-day 9-log10 kill result** | **BI test location** |
| Sentinel® XL CBRN PAPR hood and hose | Yes | Passed | Inside top surface; under hose flap |
| Sentinel® PAPR hood head cover | No | Passed | Inside top surface |
| Sentinel® PAPR hood head cover | Yes | Passed | Chin strap; inside top surface |
| Sentinel® PAPR breathing tube for use with Sentinel XL HP PAPR | No | Failed | On front end of hose; on back end of hose |
| Sentinel® PAPR breathing tube for use with Sentinel XL HP PAPR | Yes | Passed | Inside tube |
| Bullard RT Series PAPR hood | No | Passed | Inside top surface; outside top surface |
| DuPont Tyvek® 400 coverall | Yes | Passed | Outside chest; outside surface of axilla; inside surface of axilla |
| Fisherbrand face shield | No | Failed | Embedded in foam |
| Fisherbrand face shield | Yes | Failed (note: passed at 4-log10 kill) | Embedded in foam |
| 3D printed face shield (BWH/PanFab) | No | Passed | On Shield |
| iPad and case | No | Passed | On device; on case |
| N95 masks | No | Passed | Located on racks 1 and 2 on top middle, and bottom shelves |