**Supplementary Table 2**: Results of ambient particulate matter filtration efficiency evaluation performed at MIT, recorded within 24 hours of treatment. Values are an average of four upstream and downstream measurements, with standard deviation (SD) shown.

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Model** | **Cycles** | **Filtration Efficiency (SD)** | **ΔP (pa)** | **Vair (m/s)** | **T (°C)** | **Relative Humidity (%)** |
| **0.3μm** | **0.5μm** | **1.0μm** |
| **3M 1860** | 0 (control) | 97.66% (0.18) | 97.66% (0.18) | 97.66% (0.18) | 186.1 | 0.3 | 24.1 | 22.5 |
| 0 (control) | 97.53% (0.18) | 97.53% (0.18) | 97.53% (0.18) | 181.1 | 0.4 | 23.2 | 26.2 |
| 1 | 99.20% (0.08) | 99.20% (0.08) | 99.20% (0.08) | 191.5 | 0.024-0.1 | 24.3 | 22.0 |
| 2 | 98.98% (0.10) | 98.98% (0.10) | 98.98% (0.10) | 192.0 | 0.024-0.1 | 24.8 | 21.4 |
| 2 | 99.42% (0.05) | 99.42% (0.05) | 99.42% (0.05) | 194.0 | 0.024-0.1 | 24.8 | 21.4 |
| 3 | 99.36% (0.12) | 99.36% (0.12) | 99.36% (0.12) | 193.0 | 0.024-0.1 | 25.2 | 20.3 |
| 4 | 98.55% (0.07) | 98.55% (0.07) | 98.55% (0.07) | 192.6 | 0.3 | 25.1 | 20.2 |
| 5 | 98.76% (0.03) | 98.76% (0.03) | 98.76% (0.03) | 186.6 | 0.024-0.1 | 23.3 | 31.1 |
| 10 | 98.45% (0.15) | 98.45% (0.15) | 98.45% (0.15) | 180.2 | 0.4 | 23.8 | 25.5 |
| **KC/Halyard 46767 (duckbill)** | 0 (control) | 99.91% (0.02) | 99.95% (0.05) | 100.00% (0.00) | 190.1 | 0.024-0.1 | 24.3 | 22.7 |
| 1 | 99.83% (0.07) | 99.86% (0.17) | 100.00% (0.00) | 192.6 | 0.024-0.1 | 24.8 | 21.7 |
| 2 | 99.91% (0.02) | 99.98% (0.02) | 100.00% (0.00) | 193.2 | 0.024-0.1 | 24.9 | 22.1 |
| 3 | 99.90% (0.04) | 99.98% (0.04) | 100.00% (0.00) | 194.3 | 0.024-0.1 | 24.7 | 21.0 |
| 4 | 99.69% (0.06) | 99.80% (0.12) | 99.89% (0.24) | 194.5 | 0.024-0.1 | 24.8 | 20.5 |
| 5 | 99.89% (0.03) | 99.95% (0.07) | 100.00% (0.00) | 186.6 | 0.024-0.1 | 22.7 | 32.0 |
| 10 | 99.86% (0.07) | 99.97% (0.06) | 100.00% (0.00) | 186.4 | 0.024-0.1 | 23.8 | 25.4 |
| **Gerson 2130** | 1 | 96.06% (0.20) | 98.90% (0.10) | 99.68% (0.43) | 194.0 | 0.024-0.1 | 24.8 | 21.4 |
| 2 | 96.46% (0.19) | 99.08% (0.11) | 99.84% (0.19) | 193.3 | 0.024-0.1 | 24.7 | 21.7 |
| 3 | 95.17% (0.37) | 98.80% (0.26) | 99.65% (0.30) | 192.4 | 0.024-0.1 | 25.0 | 20.7 |
| **3M 8210** | 0 (control) | 98.09% (0.22) | 99.42% (0.24) | 99.82% (0.21) | 187.8 | 0.3 | 24.6 | 22.1 |
| 1 | 99.86% (0.04) | 99.99% (0.02) | 100.00% (0.00) | 192.0 | 0.024-0.1 | 24.8 | 21.7 |
| 2 | 99.52% (0.03) | 99.93% (0.04) | 100.00% (0.00) | 191.2 | 0.024-0.1 | 24.9 | 21.3 |
| 3 | 99.28% (0.06) | 99.88% (0.04) | 100.00% (0.00) | 191.4 | 0.3 | 25.0 | 20.3 |
| 4 | 98.90% (0.11) | 99.40% (0.10) | 100.00% (0.00) | 193.0 | 0.3 | 25.2 | 20.0 |
| 10 | 99.16% (0.15) | 99.77% (0.13) | 100.00% (0.00) | 179.4 | 0.4 | 23.8 | 25.2 |
| **3M 9210/37021** | 0 (control) | 99.75% (0.11) | 99.92% (0.11) | 100.00% (0.00) | 190.6 | 0.024-0.1 | 24.6 | 21.9 |
| 1 | 99.77% (0.16) | 99.83% (0.19) | 99.71% (0.37) | 193.8 | 0.024-0.1 | 24.8 | 21.5 |
| 2 | 99.70% (0.07) | 99.92% (0.07) | 100.00% (0.00) | 192.4 | 0.2 | 24.9 | 20.9 |
| 3 | 99.39% (0.18) | 99.86% (0.04) | 100.00% (0.00) | 192.0 | 0.024-0.1 | 25.0 | 20.6 |
| 4 | 98.68% (0.98) | 99.01% (0.92) | 99.05% (1.19) | 192.8 | 0.024-0.1 | 24.9 | 20.6 |